## Slide 1
EMPECS Medical Device Co., Ltd. All Rights reserved.
▣ App 화면 구성 요청
QR Scan  Warm-Up Screen
Display attachment guide instructions upon touching “QR Scan.“
After scrolling through the guide, touch the [Sensor Connect] button.
Subsequently, the camera screen is displayed, the sensor connects, and a countdown pop-up appears
The countdown runs from 30 minutes down to 0 minutes; once 0 minutes is reached, the pop-up disappears, and the first measurement value is displayed
Sensor Connect

## Slide 2
EMPECS Medical Device Co., Ltd. All Rights reserved.
▣ App 화면 구성 요청
Previous data view screen
Touching the button at the bottom right of the main screen adds a pop-up for the "View Previous Data" function
  Touch the bottom right button
  Display the "View Previous Data" pop-up
Refresh upon touching the button from step 1.
  Download previous records stored on the server
  The downloaded information is saved on the device and can be viewed even without an internet connection.
Selecting the button from step 2 loads the graph information recorded during the corresponding period
  The period corresponding to step 2 is data per sensor
  2025.09.25~2025.10.08 Data for C21ZS00102 (Sensor2)
  2025.09.15~2025.09.24  Data for C21ZS00101 (Sensor1)
  The data is ordered with the most recent sensor stacked at the top
  The graph shows the continuous data from the start date to the end date
  Touching the [X] button returns to the previous main screen $\rightarrow  navigates to the current sensor's measurement screen
1
2
Measurement period
sensor serial number
Previous screen
Return to the main screen upon touching the previous screen.
View Previous Data button  Select Period  View Previous Records  Touch [X] button  Return to Main Screen
2025. 9. 25 - 2025. 10. 8
2025. 10. 9 - 2025. 10. 23
Touch the icon
●
●
●
Upon touching the period

## Slide 3
EMPECS Medical Device Co., Ltd. All Rights reserved.
▣ App 화면 구성 요청
Sensor Screen – SC_03_01 Connection Status
Connecting / No sensor connected
Display only the sensor connection status information
  Connecting/Disconnected
Before change
After change

## Slide 4
EMPECS Medical Device Co., Ltd. All Rights reserved.
▣ App 화면 구성 요청
Sensor Screen – SC_05_01 Start Time
Start Time automatically updates to match the start date and time.
Delete Warm-Up Time section
  Check only the sensor start date and time
Change the screen layout to match the configuration shown after the right-side modification
OK
Before change
After change

## Slide 5
EMPECS Medical Device Co., Ltd. All Rights reserved.
▣ App 화면 구성 요청
Sensor Screen – SC_07_01 Data Share
Glucose data
User Profile
Sharing settings
Select the date you want to share
  Select a custom date instead of the 1, 7, or 30-day options.
  Ex) If only 1 day is selected, one day's worth of information is shared. If 1 to 6 days are selected, the information for the specified days is shared.
Select item in sharing range
  Glucose data: Summary, Range Distribution, and Glucose Graph in the Report.
  User Profile
Select sharing method
  Share via methods such as Email and SMS (.CSV, PDF)
Revocable anytime Delete
  This section reconfirms consent for information sharing, similar to a pop-up notification before sharing
Modification request

## Slide 6
EMPECS Medical Device Co., Ltd. All Rights reserved.
▣ App 화면 구성 요청
Alarm Screen – AR_01_02 Very Low Glucose
Display banner upon alarm
방해 금지모드 무시
Alarm On/Off
Select very low glucose range.
Select Sound/Vibration Mode.
  Sound only if 'Sound Only' is selected
  Vibration only if 'Vibration Only' is selected.
  Sound + Vibration if both are selected.
  Select notification sound
Display banner upon notification
Override Do Not Disturb Mode feature On/Off
  On: Alarm/Alert even when the phone is in Do Not Disturb Mode.
  Off: Alarm/Alert Off when in Do Not Disturb Mode.
   This section only applies the very low glucose alert.
Very Low glucose alert
Before change
After change

## Slide 7
EMPECS Medical Device Co., Ltd. All Rights reserved.
▣ App 화면 구성 요청
Alarm Screen – AR_01_03 High Glucose
0
High glucose alert
Alarm On/Off
Select high glucose range
  Graph line changes according to the set value
Select Sound/Vibration Mode.
  Sound only if 'Sound Only' is selected
  Vibration only if 'Vibration Only' is selected.
  Sound + Vibration if both are selected.
Select notification sound
Display banner upon alarm
After change
Display banner upon alarm

## Slide 8
EMPECS Medical Device Co., Ltd. All Rights reserved.
▣ App 화면 구성 요청
Alarm Screen – AR_01_04 Low Glucose
0
Low glucose alert
After change
Display banner upon alarm
Alarm On/Off
Select Low glucose range
  Graph line changes according to the set value
Select Sound/Vibration Mode.
  Sound only if 'Sound Only' is selected
  Vibration only if 'Vibration Only' is selected.
  Sound + Vibration if both are selected.
Select notification sound
Display banner upon alarm

## Slide 9
EMPECS Medical Device Co., Ltd. All Rights reserved.
▣ App 화면 구성 요청
Alarm Screen – AR_01_05 Rapid Change
0
Rapid Change Alarm
[Rapid Change Range]
Alarm On/Off
Select Rapid Change range
Select Sound/Vibration Mode.
  Sound only if 'Sound Only' is selected
  Vibration only if 'Vibration Only' is selected.
  Sound + Vibration if both are selected.
Select notification sound
Display banner upon alarm
Display banner upon alarm

## Slide 10
EMPECS Medical Device Co., Ltd. All Rights reserved.
▣ App 화면 구성 요청
Alarm Screen – AR_01_06 Signal Loss
Since this section is not for user configuration, revise it to match the post-modification version.
Delete the item(s) in the displayed section
Items in the alarm method can be selected.
  Sound + Vibration if both are selected.
Before change
After change

## Slide 11
EMPECS Medical Device Co., Ltd. All Rights reserved.
▣ App 화면 구성 요청
Alarm Screen AR_01_08 Lock Screen Alerts
Lock Screen (Example 1)
Main Screen (Example 2)
Display Activate item
Banner notification on the Lock Screen.
  Always display banner on the Lock Screen
  Display the latest glucose value and trend arrow.
Adding a widget to the Main Screen also adds an item to the Main Screen.
  Add widget to the Main Screen.
  Display graph, latest glucose value, and trend arrow.
These parts update with regular refreshes

## Slide 12
EMPECS Medical Device Co., Ltd. All Rights reserved.
▣ App 화면 구성 요청
Setting Screen – ST_01_01 User Screen
Feature to modify personal information upon selecting the User section.
  Name and email address cannot be changed, as they are displayed as login information
  Gender, date of birth, phone number, and type of diabetes can be modified when editing member information
Automatically select Region.
  Automatically select by recognizing the country of login.
Automatic language matching based on the country selected in Region.
  Language matching based on the selected country  if no match is found, the default language is English.
Apply the time display change feature upon selecting Time Format.
  24h: 24-hour format.
  12h: 12-hour format (AM/PM distinction for morning/afternoon)
E-mail
name
Touch

## Slide 13
EMPECS Medical Device Co., Ltd. All Rights reserved.
▣ App 화면 구성 요청
Setting Screen – ST_01_01 User Screen
Addition of the Inquiry section
Function of log data transmission
  If the user encounters an issue that they suspect is caused by a sensor defect during use
  Select the [Log Data Transmission] section.
    Information being transmitted: Data for troubleshooting, such as communication history between the sensor and the device
    Ex) Log transmission of information necessary to check/resolve issues such as sensor/device disconnection, sudden connection loss, or premature battery drain due to a sensor anomaly.
  Feature for transmitting log data from the sensor to our company.
2025.12.01
Transmission Completed.
Upon tapping Log Data Transmission, display the transmission date and Transmission Completed.

## Slide 14
EMPECS Medical Device Co., Ltd. All Rights reserved.
Thank You